DURRIDGE RAD7

RADLINK Installation Notes

June 1993

RADLINK installation into NVRAM

(a) The RADLINK software is installed as an extension to the

standard RAD7 operating software.  While the standard RAD7

software cannot be changed without a new EPROM program chip,

extension software can be installed and changed through

electronic signals sent through the RAD7 serial port.

Please note that RADLINK requires that the RAD7 have a

version 2.4 (or later) EPROM.  Also specific versions of

RADLINK must be compatible with the specific EPROM version.

(b) RADLINK installation requires a personal computer (PC)

and an appropriate serial data cable to connect it to the

RAD7.  The serial cable is the same "null-modem" cable

described elsewhere in the RAD7 serial port documentation,

SERIAL.DOC.

(c) To begin the installation, connect the serial data cable

from the RAD7 to one of the computer's COM ports.

(d) If you have not yet done so, copy the contents of the

RADLINK floppy disk onto your computer's hard disk.

(e) Turn on the RAD7.  Select ">Special" from the RAD7

keypad.  If this is the first software extension, the RAD7

will respond:

Not installed.

Install? Yes

Press the RAD7 "enter" key to begin the software loading.

The RAD7 display will go blank while it waits for the

program to be downloaded.

If extension software has already been installed, select

">Special SetBaud" and change the data rate to 1200 baud.

Then select ">Special S-Load".  Once again, the RAD7 display

will go blank while it waits.

(f) At the computer, type the transmit command from the DOS

prompt:

C:>TRANSMIT RAD0244.EXT COM1

where RAD0244.EXT is the name of the program file and COM1

is the name of the serial port.  Of course, if you are

connected to the COM2 port you should type COM2 instead of

COM1.

(g) The computer will give the following message:

*** TRANSMIT 1.0 *** 1/98  DURRIDGE Company

Input file RAD0244.EXT is open.

Communication port COM1 is open.

Prepare RAD7 to receive file.  Then press key to transmit.

You may press any computer key to begin the file transfer.

The computer will respond like this:

File transmit RAD0244.EXT ...

S00B00007344A456EC6B532

and so on, followed by continuing streams of letters and

numbers.  The same letters and numbers will flash across the

RAD7 display.

(h)  Allow the transfer to continue until it finishes.  You

will hear a beep.  The RAD7 display will say "OK." when it

finishes.  The computer will give the message "Transfer

finished."

(i) Turn off the RAD7 for 5 seconds to reset the

microprocessor.  When you turn it back on, the keypad should

operate normally.  The prompt character on the RAD7 menu

will change from the original ")" to ">", but otherwise

everything will seem exactly the same.

(j) Try out the ">Special Version" command to make sure the

extension program is functioning properly.  It should

respond with "Version: 0244", or the number of whatever

software version you have installed.