DURRIDGE RAD7

RADLINK REMOTE CONTROL PACKAGE

Version 0244

February 1994

1. Capabilities

The remote control package enables the DURRIDGE RAD7 to be

controlled from a distance by sending commands through the

RAD7's built-in serial port.  This package makes all the

ordinary RAD7 keypad commands available through a remote

device connected to the serial port.  Using a remote

terminal or computer you may, for example, change the RAD7's

operating parameters, start a test, stop a test, or receive

data from the RAD7.  The package permits users to implement

automated communication between the RAD7 and computing

equipment for the purpose of remote centralized data

collection or control.

2. Requirements

All newer RAD7 instruments already contain the necessary

electronic circuits to communicate through the serial port,

but may require software upgrades.  Older RAD7s (serial 270

and below) that have not been upgraded yet require the Real

Time Clock/Non-Volatile memory (RTC/NVRAM) upgrade.

Remote control software resides in non-volatile memory

(NVRAM), but does not decrease the amount of memory

available for storing radon data.  The remote control

software requires an upgraded RAD7 program chip (EPROM

version 2.4 or later), which allows software to be loaded

into the RAD7's non-volatile memory through the serial port.

Should the remote control program become scrambled or

corrupted through a software bug, the program must be

re-loaded into the RAD7.  For that purpose we are supplying

a special program for IBM compatible personal computers to

restore a scambled RAD7 or to perform upgrades.

The most simple implementation of remote control is through

a direct wire link from the RAD7 to a personal computer or

terminal.  This requires a serial null-modem cable from the

RAD7 to the serial port connector on the computer or

terminal.

The simple, direct serial port to serial port connection has

a range limited to around 50 feet (15 meters) by the

RS-232-C standard.  Options for extending this range include

RS-232 line boosters, current-loop and other types of

interface converters, short-haul modems, and leased-line

modems. You may use standard data modems to communicate over

the telephone system to one or more remote RAD7 monitors, so

that when you want to get some data or start a new run, just

"dial up" the instrument of your choice.

If you intend to use a personal computer to communicate with

and control an instrument, you will probably need some kind

of terminal emulator software for your computer.  Several

low-cost commercial software products are available for data

communication that allow you to send commands, receive

replies, and transfer data in and out of your computer using

its serial port.  ProComm and SmartTerm are products that we

are familiar with, but there are many others.  Microsoft

Windows has a built-in communication package that performs

these functions without additional software.

3. RAD7 Remote Command Format

The format of commands to the RAD7 serial port match, as

closely as possible, the format of the commands available to

the user at the RAD7 keypad.  For example, the command to

change the cycle time to 1 hour is "SETUP CYCLE 01:00".  The

command to turn off the audio beeper is "SETUP TONE OFF".

The command to send over the data from run number 3 in

comma-delimited form is "DATA COM 3".  The upper/lower case

of the command does not matter, nor does the numeric format

of numbers.  In the last example, "data com 03", "Data Com

3", and "dAtA coM 03.00" all work equally well.

You can use any of the ordinary RAD7 commands, but a few

extra commands have been added that you may also use, or in

some cases substitute for other commands.  These commands

all start with the word "SPECIAL".  One such extra command

is the command "SPECIAL STATUS" which gives information

about the current status of the instrument, just like "TEST

STATUS", but does not continue to update the information

every second; instead, it returns control to the user.

The Control-C character functions as a remote "menu" key

that can be used to interrupt certain activities and get

ready for a new command.  The RAD7 replies with a prompt,

the greater-than character ">", that tells you it is ready

for a new command.  When using a terminal emulator program,

you will always type commands at the prompt.

All commands must be followed by a carriage return

character, denoted here as <CR>.  No command will be

activated until the <CR> goes through.  Once you have typed

a command, always end with a carriage return keystroke.

This key may be marked "Enter" or "Return" on your keyboard.

If the RAD7 cannot understand your command, such as if you

typed words in the wrong order or misspelled something, it

will respond with "?ERROR", followed by a list of acceptable

command words.

4. Special Keys

If you are using a terminal emulator package, you can gain a

few special keys by making sure your terminal has been setup

to give ANSI standard escape codes for the function keys and

cursor control keys.  Set your terminal for either ANSI,

VT-52, or VT-100 mode to get these functions.  The function

keys F1, F2, F3, and F4 act as a remote RAD7 keypad,

corresponding to the RAD7 keys <menu>, <enter>, <backward>,

and <forward>.  As mentioned previously, the Control-C

character also acts as a remote <menu> key.  The

backspace/delete key on your keyboard allows you to correct

misspelled commands before the carriage return.  If this key

does not work, then Control-H may do the same function.

5. Special commands

A series of special commands have been added to the ordinary

RAD7 commands to provide functions of special interest to

remote control users:

Special Ident - Output the RAD7 identification sequence,

including firmware version, hardware model number, unit

serial number, and last calibration date.

Special SPrOn - Re-direct subsequent output from the

infra-red printer to the serial port.  In other words,

everthing that would ordinarily be printed will shoot out

the serial port, but nothing will be printed, even when you

say "Print".  One reason to use this might be to move the

data very quickly into a computer without waiting for the

(slow) infra-red printer link.  You can cancel the

re-direction order and restore the use to the infra-red

printer with the "Special SPrOff" command.  When you turn

off the RAD7 and turn it on again, it always restores output

to the infra-red printer.

Special SPrOff - Cancel the printer to serial port

re-direction, so that output can go to the printer again.

Special SetBaud - Set the serial port bit rate.  The

following standard speeds are available: 300, 600, 1200,

2400, 4800, 9600, and 19,200 bps.  The other communication

settings are always 8 bit, no parity, and 1 stop bit.  The

RAD7 remembers the serial port speed when you power down.

Note that at the highest serial speed settings the RAD7 may

not be able to keep up with incomming character strings

unless the characters are "paced".  An "echo-wait" strategy

will avoid this problem.  Also note that the RAD7 recognizes

XOFF/XON flow control protocol when sending data.

Special Status - Gives a snapshot of the RAD7 status page,

including run and cycle numbers, countdown timer, last

reading, temperature, humidity, and so on.  This is

basically the same information that you can get with "Test

Status", but it gives the data in one shot and does not

continue to update every second.

Special Start - Same as "Test Start", but does not go into a

continuously updating status display.

Special Stop - Same as "Test Stop".

Special S-Load - Used to load special software into the RAD7

through the serial port.

Special Version - Output the special extension version

number.

Special Model - Output the RAD7 hardware version number.

Special Serial - Output the RAD7 unit serial number.

Special Beep - The RAD7 gives an audible beep tone.  Does

not make any sound if the tone setting is "Off".

Special ComAll - (Version 0244/940221 or later) - Output

complete set of RAD7 test data (up to 1000 readings) to the

serial port in comma delimited format.

Special SPrAll - (Version 0244/940221 or later) - Output

complete set of RAD7 test data (up to 1000 readings) to the

serial port in standard, readable 24 column printer format,

without affecting the infra-red printer.