DURRIDGE RAD7

RS-232 SERIAL PORT

Version 2.4

June 1993

1. Capabilities

The RAD7's built-in serial port allows you to transfer

data to your personal computer.  Once the data has been

transferred, you may use the software of your choice to

manage the data and create tables, graphs and reports.

The standard RAD7 firmware enables one-way communication

from the RAD7 to another device.  Two-way communication

may be added through the RADLINK remote control package.

The RAD7 serial port follows RS-232 convention for

compatibility with nearly all personal computers,

including the IBM-PC, XT, and AT, as well as their

"clones".

In order to take advantage of the RAD7 data transfer

capability, you will need appropriate hardware (PC

serial port and serial cable), as well as appropriate

software (data capture), which will be explained below.

2. Hardware bridge

Most PC's come with one or more serial port connectors,

which can connect the computer to a modem, mouse,

printer, or another device.  The connectors are usually

25-pin male "D-sub" connectors on IBM-PC and XT

computers, and 9-pin male D-sub connectors on AT, '386,

and '486 computers.

If your computer does not already have such a connector

available, you can probably purchase a "serial adapter

card" and have it installed for under $100.

You will need the proper cable to link the RAD7 serial

port to your computer's serial port.  The cable should

be a "null-modem" type, as the RAD7 serial port emulates

a computer rather than a modem.  If you prefer, you can

use a normal "modem" cable along with a "null-modem"

adapter on one end or the other.  If your computer has a

25 pin connector, you will need a 25-pin to 9-pin

null-modem or "XT to AT" cable.  If your computer has a

9-pin connector, you will need a 9-pin to 9-pin or "AT

to AT" cable.  You can purchase the necessary cable or

cable/adapter combination at your local computer store

or Radio-Shack store, or order it from DURRIDGE.

Apple Computer products generally do not have standard

RS-232 serial port connectors, so you may need a special

adapter to connect to the RAD7.  If you are interested

in interfacing your RAD7 to an Apple Macintosh, for

instance, call DURRIDGE for assistance.

3. Software bridge

In order to receive data on your computer, you will need

software that can "capture" the data and store it to

memory or disk file.  You have several choices.  Most

terminal emulator programs, such as PROCOMM or

SMART-TERM, allow you to capture serial data to a disk

file.  Microsoft Windows has a built-in terminal

emulator package with a data capture function (they call

it "Receive Text").  These programs are just fine, but

may be somewhat cumbersome to use for a novice.  If you

do not already use a terminal emulator for modem

communications, I do not recommend your running out and

buying one.  You have another choice.

DURRIDGE has a simple-to-use cut-down capture program for

the PC/XT/AT series.  CAPTURE 1.2 takes data from the

serial port and passes it directly to a disk file.

Since the program does only one thing, it is very simple

to use.  CAPTURE 1.2 requires a PC/XT/AT series or

compatible computer that runs some version of DOS.

4. Applications

We have successfully transferred RAD7 data to several

spreadsheet programs, including Lotus 123 and

Quattro-Pro.  Once you capture the data and put into a

file, the spreadsheet program can easily pick up the

data and incorporate it into your worksheet.  The

spreadsheet program has a special command to take in

external data.  In Lotus 123, for instance, you use the

"/ FILE IMPORT NUMBERS" command to enter the data into

your worksheet.  Once the data has been entered, you

can use the 123 program to calculate averages, or

prepare tables, charts, and graphs.

Database programs, such as DBASE and Q&A, also allow you

to import RAD7 data that you have captured.  Most of

these programs allow you to create tables, graphs and

reports as well.

5. Technical specifications for the RAD7 serial port

The RAD7 serial port follows RS-232C convention for

signal levels.  Positive voltage (+3V to +15V) indicates

logic state 0 (SPACE), while negative voltage (-3V to

-15V) indicates logic state 1 (MARK).

The connector pin-out follows the IBM PC-AT convention

for the 9 pin serial port.  The handshaking lines (DTR,

DSR, RTS, and CTS) are not fully implemented in

software, and should be considered non-functional.

RAD7 Serial Port Pin Assignments (same as IBM PC-AT)

Pin     Name                        Comment

1       Carrier Detect (CD)         Not functional

2       Receive Data (RD)

3       Transmit Data (TD)

4       Data Terminal Ready (DTR)   Not fully functional

5       Signal Ground (SG)

6       Data Set Ready (DSR)        Not fully functional

7       Request To Send (RTS)       Not fully functional

8       Clear To Send (CTS)         Not fully functional

9       Ring Indicator (RI)         Not functional

The RAD7 serial port has complete hardware

implementation for two-way communication at 300, 600,

1200, 2400, 4800, 9600, and 19,200 bits per second

(bps); these speeds are available through the RADLINK

remote control package.  The default speed is 1200 bps.

RAD7 Communication Parameters

Default rate    1200 bps

Data bits       8 bits

Parity bit      None

Stop bits       1 bit

6. Serial Port Output

All stored test data are available for serial port

output through the "> Data Com ##" command.  Specify the

run number in the command line.  Output from the RAD7

serial port is entirely ASCII numeric. Each cycle

produces a record containing 23 fields.  Carriage-return

line-feeds separate the records.  Within each record,

commas separate the fields.  Fields may have leading

zeros, extra space characters, trailing decimals, etc.,

but these will not bother most data-base or spread-sheet

programs.

RAD7 Data Record - 23 data fields per record:

1.      Test number - ranges from 001 to 999

2.      Year

3.      Month

4.      Day of month

5.      Hour

6.      Minute

7.      Total counts

8.      Live time - (in minutes, resolution of 0.1 min.)

9.      Percent of total counts in window A (Res. 0.1%)

10.     Percent of total counts in window B

11.     Percent of total counts in window C

12.     Percent of total counts in window D

(Note: These 4 windows will not add up to 100%

since counts can come in to channels below or

above these windows.)

13.     High voltage level - typically 2200 V to 2300 V

14.     High voltage duty cycle - Ranges 0 to 150%, with

typical values of 10-20%

15.     Temperature - may be in degrees Celsius or

Farenheit, depending on RAD7 setup

16.     Relative humidity of sampled air - should be

kept below 10% for most accurate test

17.     Leakage current - range 0-255, above 50 is cause

for concern

18.     Battery voltage - below 6.00V is a discharged

battery.  Does not affect accuracy of test, but

indicates need for recharge.

19.     Pump current - Ranges 0-250 in units of mA.

Typically 30-40mA, above 50mA indicates

possible clogged filter or obstruction.

20.     Flags byte - individual bits indicate settings;

Bit0 set to indicate pump is in TIMED mode

Bit1 set to indicate pump is ON continuously

Bit2 not defined

Bit3 set to indicate tone is in GEIGER mode

Bit4 set to indicate beeper is activated

Bit5 set to indicate that spectrum will print

after each test

Bit6 set for multiple (recycle) test mode, else

stops after test is complete

Bit7 set for SNIFFER test mode. Clear is NORMAL.

21.     Radon concentration result in pCi/l, Bq/m3, cpm,

or # counts, depending on the units the RAD7

has been setup to use.

22.     Uncertainty - 2 sigma error of concentration in

the SAME units as the concentration (not %).

23.     Units byte - Individual bits must be extracted

to get settings;

Bits 0 and 1 indicate concentration unit

00 = cpm

01 = # of counts

10 = Bq/m3

11 = pCi/L

Bit2 through Bit6 are not defined

Bit7 indicates temperature unit (0 = Farenheit,

1 = Celsius)

A run can contain many cycles of data, with one record

per cycle.  Here is an example of a single record:

009,91,10,29,04,18,4823.,337.8,45.4, 2.9,46.6,0.3, 2201,

14, 23.7, 5, 7,7.09, 00,125,28.32743,.8500846,255<CRLF>