THORON MEASUREMENT WITH THE RAD7

Version 2.5

May 1994

THORON AND RADON

Thoron is an isotope of the element radon having an

atomic mass of 220, so it is also known as radon-220.

The word "radon" without a mass number almost always

refers to radon-222.  Thoron and radon have very similar

properties.  They are both chemically inert radioactive

gases that occur naturally from the decay of radioactive

elements in soils and minerals.  Both thoron and radon

are members of decay chains, or long sequences of

radioactive decay.

While radon results from the decay of natural uranium,

thoron results from the decay of natural thorium.  Both

uranium and thorium are commonly found in soils and

minerals, sometimes separately, sometimes together.  The

radioactive gases radon and thoron that are produced in

these soils and minerals can diffuse out of the material

and travel long distances before they themselves decay.

Both radon and thoron decay into radioactive decay

products, or progeny, of polonium, lead, and bismuth

before finally reaching stable forms as lead.

Thoron and radon and their respective progeny differ

very significantly in their half-lives and in the

energies of their radiations.  While radon has a

half-life of nearly 4 days, thoron has a half-life of

only 55 seconds.  Since thoron is so short lived, it

cannot travel as far from its source as radon can before

it decays.  It is commonly observed that compared to

that of radon gas, a much smaller fraction of the thoron

gas in soil ever reaches the interior of a building.

Even so, thoron can still be a hazard since its progeny

include lead-212 which has a half-life of 10.6 hours,

more than long enough to accumulate to significant

levels in breathable air.

THORON MEASUREMENT ISSUES

Many difficulties impede the accurate measurement of

thoron gas.  The presence of radon gas (often found

together with thoron) can interfere with a measurement.

The short half-life of thoron gas makes some aspects of

the measurement easier, but makes sampling method a

critical issue.  Thoron concentration can vary greatly

through a space, depending on the speed and direction of

air movement as well as turbulence.  The position of the

sample intake can strongly affect the results.

For many instruments, radon and thoron interfere with

each other.  Generally speaking, it is difficult to

measure one isotope accurately in the presence of the

other.  But compared to other instruments, the RAD7 is

much less susceptible to radon-thoron interference due

to its ability to distinguish the isotopes by their

unique alpha particle energies.  The RAD7 separates

radon and thoron signals and counts the two isotopes at

the same time with little interference from one to the

other.

Some issues of concern in measuring radon do not apply

to thoron.  The short half-lives of thoron (55 seconds)

and its first decay product (Po-216 - 0.15 seconds) mean

that thoron measurements can be made quickly and in

rapid succession, since there is little concern with

growth and decay delays.  The RAD7 responds virtually

instantly to the presence of thoron; its time constant

for response to thoron is less than 1 minute.  The chief

limit on the thoron response speed is the pump's ability

to fill the internal cell.  And the RAD7 clears just as

rapidly when you purge the instrument with thoron-free

air.  In fact, you need not purge the instrument at all

as thoron's short half-life ensures that it will be gone

in a few minutes.

In thoron measurement the sample pump must run in a

continuous fashion, at a steady consistent flow rate.

If the flow rate of the sampling pump changes, then the

RAD7 thoron result will also change.  Flow rate affects

the amount of thoron in the internal cell, since a

significant fraction of the thoron decays in the sample

intake system as well as within the instrument.  For the

most accurate thoron measurements we recommend that you

use a consistent sample intake system (always use the

same hose and filter arrangement) and pay special

attention to air flow rate.  Be sure the filters, hoses,

and RAD7 inlet and outlet ports remain free from

obstruction.  Use a flow meter (rotameter or "floating

ball" type) to check that the flow remains consistent.

Note that with the RAD7 the flow rate affects thoron,

but does not affect radon due to its much longer

half-life.

The RAD7 can only measure thoron concentrations of the

air at the point of sample intake.  Since thoron varies

from place to place depending on the motion of the air,

the instrument operator may find it necessary to make

measurements in several locations to properly assess a

thoron situation.  Fortunately, rapid-fire thoron

measurements are very easy to do with the RAD7.

CALCULATION AND INTERFERENCE CORRECTION

The RAD7 calculates thoron concentration on the basis of

the count rate in spectrum window B which is centered on

the 6.78 MeV alpha line of Po-216, the first decay

product of thoron gas.  To further avoid interference

from radon, the RAD7 applies a correction to the thoron

count rate to compensate for a small percentage of

"spillover" from window C.

As of version 2.5, RAD7 calculates radon concentration

from the count rate in window A (Sniff mode) or windows

A plus C (Normal mode).  The RAD7 compensates for

interference from the long-lived progeny of thoron (10.6

hours) by applying a correction to the radon count rate

in both Sniff and Normal modes.  The correction is based

on a fixed fraction of the count rate in the D window

(around the 8.78 MeV peak of Po-212) which predicts the

amount of thoron progeny activity in the A window (due

to the 6.05 and 6.09 MeV peaks of Bi-212).

Note that the uncertainty figures given with each

reading include the effect of these corrections.

AVOIDING LONGER LIVED DECAY PRODUCTS

Although the RAD7 now corrects for the buildup of the

long-lived thoron progeny (10.6 hour), we recommend that

you avoid unnecessary exposure of the instrument to high

levels of thoron for long periods of time.  The presence

of these long-lived progeny can make low level radon

measurements somewhat less accurate than would otherwise

be possible.  But if you err, the 10.6 hour half-life of

the thoron progeny makes for a temporary inconvenience

of a few days at worst.

DECAY CORRECTION, FLOW RATE, AND THORON CALIBRATION

As discussed above, thoron's rapid decay causes the

intake path and the air flow rate become important

factors in calibration.  The RAD7 factory calibration

for thoron is based on a standard RAD7 inlet filter, a

standard 3-foot long, 3/16 inch inner diameter vinyl

hose, and a standard small (6 inch) drying tube.

Deviation from this arrangement can change your thoron

results.  For example, if you were to use a very long

hose for thoron sampling, then the sample might decay

significantly before it ever reached the instrument

inlet.  The same thing might happen if you substituted

the small drying tube with the large drying column.  If

you were to use a non-recommended inlet filter, the flow

might be restricted enough to greatly lower the result.

CALCULATING SAMPLE DECAY

The thoron concentration at the inlet of the RAD7, C1,

can be expressed mathematically as

C1 = C0 * exp( -L * V1 / q )

where C0 is the original sample concentration, V1 is the

volume of the sample tube + drying tube + filter (around

50 mL), q is the flow rate (around 650 mL/min), and L is

the decay constant for thoron (.756 /min).  A typical

value for C1/C0 is then

C1/C0 = exp( -.756 * 50 / 650 ) = .943 = 94.3%

This is the number NITON assumes in the factory

calibration.  Adding a few extra feet of hose will not

matter much (about 0.5% per foot), but if we were to use

a 100 foot hose instead (V1 is around 580 mL) then the

same calculation would give .509 or 50.9%, a significant

reduction from 94.3%!

CALCULATING INTERNAL CELL CONCENTRATION

The sample decays slightly in going from the RAD7 inlet

to the internal cell, due to internal hose and filter

volumes.  This decay can be calculated in a similar

fashion to the above, giving the internal cell inlet

concentration, C2, about 95.5% of C1.  Within the RAD7

internal cell, the equilibrium thoron concentration, C3,

will be determined by the following formula:

C3 = C2 / ( 1 + L * V2 / q )

where L and q are as above, and V2 is the volume of the

internal cell (around 750 mL).  Typical values then give

C3/C2 as

C3/C2 = 1 / ( 1 + .756 * 750 / 650 ) = .534 = 53.4%

Multiplying this result by the sample decay factors

calculated above, we obtain an overall concentration in

the internal cell of 48.1% of the original sample.

Recognising the uncertainty of several of the inputs to

these formulas, particularly the flow rate, we will

round the overall result to 50%.

INTERNAL CELL THORON SENSITIVITY CALIBRATION

Preliminary investigations at NITON have shown that the

RAD7's internal cell thoron sensitivity in cpm/(pCi/L)

is identical to its radon Sniff mode sensitivity, to

within 25%.  We have no reason to expect any sizeable

difference between the thoron and radon Sniff mode

sensitivity values, so we are presently assuming that

the two values are indeed identical, and stating a

thoron calibration precision of +/- 25%.

When an opportunity to calibrate against an absolute

thoron standard becomes available, we will be able to

state the thoron sensitivity with much higher certainty.

For now we are estimating our overall thoron sensitivty

to be 50% of the radon Sniff sensitivity to account for

sample decay in the intake and internal cell.  The RAD7

has a typical radon Sniff sensitivity of .2 cpm/(pCi/L),

so we estimate the typical thoron sensitivity to be

around .1 cpm/(pCi/L).

SETTING UP A THORON MEASUREMENT

Sniffing for thoron is much the same as sniffing for

radon, except it tends to be a little faster.  If you

are just "prospecting", you probably will not be very

interested in getting the most accuruate results

possible, so technique is not critical.  But if you are

trying to make an accurate measurement, technique is of

great importance.

For accurate thoron measurement, always use the same

sample taking arrangement.  Keep the sample tubing

short: no more than 6 feet (2 meters) total length.  Use

one of the small drying tubes supplied with the

instrument, filled with fresh (blue) desiccant, of

course.  Always use an inlet filter, free from flow

restrictions or clogs.  Avoid obstructing the intake of

the sample tube.  For the most accurate results, check

the flow rate with a flow gauge to be sure it is

consistent from measurement to measurement.

Use the RAD7's Setup Protocol command to choose the

Thoron protocol for a 5 minute repeating cycle.  Be sure

the instrument has been "dried out" before making a

measurement.  Position the sample tube intake and start

the test.

THORON MODE

Thoron mode causes the RAD7 to print both thoron and

radon concentrations (in pCi/L or Bq/m3) in continuous

data logging or in subsequent printing of data.  Thoron

mode also directs the automatic pump setting to

continuous pump operation to assure a fresh sample.

The setup parameter "Setup Thoron" allows you to select

Thoron mode On/Off, for configuring a particular test to

perform thoron readings.

The protocol "Protocol: Thoron" provides a standard test

for sniffing both radon and thoron in 5 minute cycles.