DURRIDGE RAD7

IMPROVEMENTS TO OPERATING SOFTWARE

Version 2.5

NITON Corporation has made a number of improvements to the RAD7

operating software.  Some of these improvements are not

in the latest version of the RAD7 Operation Manual, so

this brief description will serve as temporary documentation.

NEW FEATURES SINCE VERSION 2.0

2.1  Real-Time Clock Calendar support

Non-Volatile RAM data storage support

Quick save-and-restart function

2.2  Automatic mode switching

Standard deviation calculation

Key pad lock-out feature

Setup review

2.3  Grab sampling protocol

Water testing capability

2.4  Special software extension capability

2.5  Changes in protocol interface

Thoron mode and thoron calculation

Changes to radon calculation

REAL-TIME CLOCK CALENDAR SUPPORT

The Real-Time Clock Calendar (RTC) upgrade provides an

independently powered clock circuit that will maintain

time-of-day and calendar date for as long as 10 years,

regardless of whether the RAD7 is on or off.  The RTC

circuit has a specified accuracy of better than 1 minute

per month at room temperature.  Extended operation or

storage    at higher temperatures may affect the

accuracy of the the clock, and shorten the lifetime of

the power source.

Operation of the instrument is essentially identical to

before the upgrade, except there will be no need to set

the time-of-day and date upon power-up.  In case you

wish to adjustment the time-of-day, or to change time

zones or enter daylight savings time, use the ">Setup

Clock" command, as before.

NON-VOLATILE RAM DATA STORAGE SUPPORT

The Non-Volatile RAM (NVRAM) upgrade provides over

32,000 bytes of non-volatile memory to allow storage of

up to 1000 complete radon readings with or without the

RAD7 on.  The NVRAM circuit uses the same power source

as the RTC circuit, so data will be retained for up to

10 years at room temperature.

With the RTC and NVRAM upgrades, RAD7 users will no

longer feel need to keep the RAD7 on when not being

used.  This will help the user to avoid accidentally

running down the main batteries by forgetting to plug in

the power cord.  Simply turn the power switch off to

avoid draining the battery, and the data will still be

available the next time you power the instrument.

Several commands are available to help the user manage

the stored data.  More information can be found in the

manual.

QUICK SAVE-AND-RESTART FUNCTION

This function allows the user to end a sniff test, store

it to memory, and start up a new sniff test, all using a

single key.  This function only operates from the sniff

mode.

From the status display (showing countdown timer), press

the enter key once.  The display will show "Save and

restart? Yes".  Press the enter key once more to confirm

your intention.  If you do not wish to save-and-restart,

press one of the arrow keys so that the word "No"

appears rather than "Yes".  Press the enter key to

select "No" and the status display will reappear, with

the countdown continuing as before.

AUTOMATIC MODE SWITCHING

This feature includes an additional mode setting along

with Sniff and Normal.  From ">Setup Mode" you may now

select "Auto" which starts a test run in Sniff mode, but

switches over to Normal after the run has continued for

3 hours.  This allows time for equilibrium of the longer

lived radon daughter isotopes.  The earliest part of the

run will have the benefit of the Sniff mode's quick

response, while the later parts of the run will have the

benefit of the Normal mode's statistical precision.

We recommend that you use the Auto setting for screening

tests and all other tests that attempt to measure the

average concentration over a 24 or 48 hour period.

Using the Auto setting, you will no longer be required

to "throw away" the first three hours of data, or

calculate adjustments to correct for disequilibrium.

The mean concentration reported in the run summary

should accurately reflect the actual mean.

As always, for the most accurate results be sure to

purge the RAD7 with desiccated air before beginning the

test.  Be careful to avoid short-lived background due to

previously sampled air.

STANDARD DEVIATION CALCULATION

The run summary now includes standard deviation (S.D.)

of the run's concentrations, along with the mean, high,

and low concentrations.

The standard deviation is calculated by the usual

formula, using n - 1 as the population parameter.  Note

that no attempt has been made to correct for the

statistical precision of the individual data points, or

to adjust for differing test periods.

Standard deviation (S.D.) is reported in the same unit

as concentration; to get S.D. in percent, divide by the

mean concentration and multiply by 100.

KEY PAD LOCK-OUT

The ">Test Lock" command allows the user to lock the key

pad to prevent casual access to RAD7 commands or data.

Once the command has been invoked, the display will show

"** NITON RAD7 **".  Any test underway will continue

until completion, and printing and data logging will

operate as usual.  Pressing the keys individually will

have no effect on the operation of the instrument, or in

any way allow information to be displayed.

If someone should reset the instrument, by toggling the

on/off switch, the keypad will remain locked.  An

uncompleted test, however, will not be continued.

To unlock the key pad, hold down the enter, forward, and

backward arrow keys simultaneously for 3 to 5 seconds,

or until you hear the RAD7 beep.  Release the three keys

and then quickly press the menu key.  If you do not

press menu within 2 seconds the key pad will remain

locked.  If the RAD7 does not unlock on the first

attempt, try different three key hold-down times from 2

to 6 seconds.

SETUP REVIEW

The ">Setup Review" command allows you to display and

print a listing of the current instrument settings,

including CycleTime, Recycle, Mode, Pump, Tone,

PrintForm, and Unit.  This allows you to check that you

have properly setup the instrument for your intended

use.

Also, the instrument will automatically print its

settings when you first turn the power on.

GRAB SAMPLING

The "Grab" pump setting, available through the ">Setup

Pump Grab"  command or through several new Protocols,

initiates a standard grab sampling sequence at the

beginning of a run.  When you start a new run with

"Grab", the pump will run for exactly 5 minutes,

followed by a 5 minute equilibrium delay, before the

sample counting period begins.  The pump does not run at

all during the counting period.  As usual, you can

determine the length of the count period by multiplying

the cycle time by the recycle number.  The total time

required to complete a test is the pump sample time (5

minutes) plus the delay period (5 minutes) plus the

count time.

The "Grab" protocol, available through the ">Setup

Protocol Grab" command, sets the pump to "Grab", the

mode to "Sniff", the cycle time to 5 minutes, and the

recycle number to 4.  A complete test under the Grab

protocol takes 30 minutes, with 5 minutes for the pump

sample, 5 minutes for the delay, and 20 minutes for the

counting period.

WATER TESTING

We have added two new options to the Mode setting to

permit the use of the RAD-H2O radon in water apparatus.

The "Wat-40" setting is intended for use with the 40 mL

water sample vials, and the "Wat250" setting is intended

for use with the 250 mL water sample vials.

We have also added two new protocols under the ">Setup

Protocol" command.  "Wat-40" and "Wat250" are intended

for use with the RAD-H2O apparatus to perform a complete

water analysis in 30 minutes.  The water test sequence

is similar to the Grab protocol, with 5 minutes for the

pump, 5 minutes delay, and 20 minutes counting period.

SPECIAL SOFTWARE EXTENSION

The newest versions of RAD7 firmware allow the

installation of special application software through the

serial port.  At present, this capability is used to

install the RADLINK remote control software into

non-volatile RAM memory, and provides a place for new

commands to be added to the RAD7 menu under ">Special".

A RAD7 that does not have extension software will state

"Not installed"  when the user selects ">Special".  It

will then ask the user whether or not he wishes to

install extension software.  After answering "yes", the

user must initiate the transfer of the extension

software from a PC to the RAD7 serial port.  After the

software installation finishes, ">Special" will head the

list of new commands.

The special extension capability permits NITON to

accommodate our customer's requests for customized

instrument operation.  If you have an application for

the RAD7 that cannot be implemented without modification

of the RAD7 operation, contact Stephen Shefsky at NITON

Corporation to discuss the possibility of custom

software extensions.

CHANGES IN PROTOCOL INTERFACE

The Setup Protocol command now operates more like the

other Setup commands by indicating the active protocol,

or the last protocol to be selected.  If any critical

setup parameters (Cycle, Recycle, Mode, Pump) have been

changed since the last protocol selection, then the

display will indicate "(None)" to emphasize that the

setup no longer matches one of the standard protocols.

The Setup Review now indicates the active protocol along

with the various setup parameters.  Again, if any of the

critical parameters have been changed, the protocol will

be "(None)".

THORON MODE AND CALCULATION

The RAD7 now allows simultaneous indication of radon and

thoron concentrations in pCi/L or Bq/m3.  Thoron mode

causes the RAD7 to print thoron concentrations along

with the radon concentrations in continuous data logging

or in subsequent printing of data.  Thoron mode also

directs the automatic pump setting to continuous pump

operation to assure a continuously fresh sample.

The "Pump: Timed" setting has been renamed as "Pump:

Auto" to indicate its new influence in Thoron mode.

When Thoron mode is not selected the pump behaves in

Auto mode exactly as it did before in Timed mode.

A new setup parameter (Setup Thoron) allows you to

select Thoron mode On/Off, and a new Protocol (Setup

Protocol Thoron) provides for simultaneous sniffing for

radon and thoron in 5 minute cycles.

The calculation of thoron concentration is based on the

count rate in the RAD7 spectrum window B which is

centered around the Po-216 alpha line at 6.78 MeV.  To

avoid interference from radon, the RAD7 applies a small

correction to the thoron count rate to compensate for

"spillover" from window C.

For more information about thoron measurement consult

the NITON document "Thoron Measurement with the RAD7".

CHANGES TO RADON CALCULATION

To make the radon measurement less prone to interference

from thoron in the normal mode, the RAD7 now calculates

normal mode count rate from the sum of the count rates

in the A and C windows only rather than the sum of the

A, B, and C windows.

In addition, the RAD7 compensates for interference from

the long-lived progeny of thoron by applying a

correction to the radon count rate in both Sniff and

Normal modes.  The correction is based on a fixed

fraction of the count rate in the D window (around the

8.78 MeV peak of Po-212) which predicts the amount of

thoron progeny activity in the A window (due to the 6.05

and 6.09 MeV peaks of Bi-212).

The effect of these changes on thoron-free radon

measurement is small, but does require small adjustments

to the RAD7 calibration paramenters.  For that reason we

suggest that the instrument be re-calibrated when

upgrading to version 2.5.